# Changing The Machine ID in GeoStar

This document explains how to change the Machine ID in GeoStar.

### When to Follow This Procedure

- During installation to ensure the Machine ID in the GeoStar software matches the instrument serial number printed on the instrument chassis.

- When it is discovered that the Machine ID does not match the instrument serial number, typically due to an omission during assembly or installation.

### What the Procedure Requires

- GeoStar.

- About five minutes of time.

# The Procedure

Changing the Machine ID within GeoStar will result in the creation of a new system configuration. Once that process is started then the software will have to be fully configured before it can be used.

To make the change faster and less prone to configuration mistakes, this procedure outlines how to edit the configuration files directly.

Close GeoStar.

The Machine ID is stored across a number of configuration files. The following text identifies the files and highlights the changes that are required.

Note: All files reside in the GeoStar ./conf folder and paths are given relative to this location. Consult the GeoStar documentation if you require assistance finding the ./conf folder.

Note: For the purpose of demonstration, the procedure will show edits to change an existing Machine ID value of “MID0001” to a new value of “MID0002”. Real Machine ID values will have a different format.

## .\config\system.xml

Edit the file. The file contains a parameter for the Machine ID.

<MachineID>MID0001</MachineID>

Change the old Machine ID to the new value:

<MachineID>MID0002</MachineID>

Save and close the file.

## .\config\io_config_vMID0001.xml

This is the configuration file for the IO channels. The Machine ID is present in both the filename and the file contents.

Rename the file to contain the new Machine ID. In this example the file would be renamed to: .\config\io_config_vMID0002.xml

Edit the file. The file contains a parameter on the second line:

<io version="MID0001">

Change the old Machine ID to the new value:

<io version="MID0002">

Save and close the file.

## .\config\gas_config_vMID0001.xml

This is the configuration file for the gas control panel. The Machine ID is present in both the filename and the file contents.

Rename the file to contain the new Machine ID. In this example the file would be renamed to: .\config\gas_config_vMID0002.xml

Edit the file. The file contains a parameter on the second line:

<gasconfig version="MID0001">

Change the old Machine ID to the new value:

<gasconfig version="MID0002">

Save and close the file.

## .\config\gas_config_vMID0001.png

This file contains the background image for the gas control panel. The Machine ID is present in the filename.

Rename the file to contain the new Machine ID. In this example the file would be renamed to: .\config\gas_config_vMID0002.png

## Registration and Testing

Any mistakes made during file editing will result in GeoStar crashing on start up, or, reverting to offline mode. If GeoStar fails to start carefully check all configuration files for typing mistakes.

After making these changes, always start GeoStar and test that everything operates as expected.

On start up GeoStar will ask for a new registration key. Follow the prompt to the website and enter the installation details to receive a new registration key.

[End of Guide]